RDPDZ-104T  Comprehensive Partial Discharge (PD) Inspection Instrument

1.Product Overview

Partial discharge (PD) is a pulsed discharge that generates a series of physical phenomena (such as light, sound, electrical signals, and mechanical vibrations) and chemical changes inside power equipment and the surrounding space. These accompanying physical and chemical changes provide detectable signals for monitoring the internal insulation condition of power equipment. When insulation defects occur inside high-voltage electrical equipment, PD signals are inevitably generated. Through the detection and analysis of PD signals, potential insulation hazards in high-voltage electrical equipment can be identified, preventing the further expansion of potential accidents.

Developed by our company, this multi-functional handheld PD detector adopts detection methods including Transient Earth Voltage (TEV), Acoustic Emission (AE), Ultra-High Frequency (UHF), and High-Frequency Current Transformer (HFCT) to test PD in equipment. It can measure PD magnitude and display waveform patterns, with functions for storing and retrieving 2D/3D 图谱. The instrument effectively evaluates PD conditions in electrical equipment and is suitable for PD detection of GIS, switchgears, transformers, power cables, and other electrical equipment. The PD severity of electrical equipment is assessed through the following indicators:

PD Intensity Detection: Measures PD signals within one power frequency cycle and characterizes PD intensity by the maximum value (in dB) in the discharge pulse sequence.

PD Frequency Detection: Measures PD signals within one power frequency cycle, extracts discharge pulses, and characterizes PD frequency by the number of pulses.

The instrument is portable and easy to operate, with all detection processes having no impact on the operation of high-voltage equipment. It supports multi-cycle observation of measured signals, frequency identification of discharges, and analysis through multiple modes, enabling clear fault diagnosis.

2. Product Features

☀Configurable with different sensors to realize PD detection for almost all types of high-voltage electrical equipment.

☀Provides multiple discharge patterns including time-domain wave forms, PRPD (Phase Resolved Partial Discharge), and PRPS (Phase Resolved Pulse Sequence) for analyzing different discharge types.

☀ User-friendly human-machine interface (HMI) facilitates data management for different equipment.

☀ Built-in AE sensor and TEV sensor, with external interfaces for dedicated sensors of transformers, GIS, overhead lines, cables, etc.

☀ Adopts non-intrusive detection method, no power outage required during testing, no additional high-voltage source needed, more convenient to use than traditional pulsed PD detectors.

☀ Test bandwidth range: 30kHz ~ 2.0GHz, suitable for various detection principles in different frequency bands.

3. Parameter Specifications

| Measurement Category | Parameter | Specifications |
| --- | --- | --- |
| TEV Measurement | Measurement Range | 0dB ~ 60dB |
|  | Resolution | 0.1dB |
|  | Accuracy | ±1dB |
|  | Maximum Pulses per Cycle | 1400 |
|  | Measurement Bandwidth | 3 ~ 100MHz |
| AE Measurement | Measurement Range | 0dB ~ 60dB |
|  | Resolution | 0.1dB |
|  | Accuracy | ±1dB |
|  | Frequency Range | 20 ~ 200 kHz |
| AA, AA1, AA2 Measurement | Measurement Range | 0dB ~ 60dB |
|  | Resolution | 0.1dB |
|  | Accuracy | ±1dB |
|  | Sensor Center Frequency | 40 kHz |
| UHF Measurement | Detection Bandwidth | 300 ~ 2000MHz |
|  | Measurement Range | 0dB ~ 60dB |
|  | Sensitivity | ＜1dB |
|  | Sensor Bandwidth | 300 ~ 2000MHz |
| HFCT Measurement | Sensor Transfer Impedance | 5mV/mA |
|  | Detection Frequency | 1 ~ 30MHz |
|  | Resolution | 0.1dB |
|  | Sensitivity | ≤50pC |
| Hardware | Housing Material | ABS |
|  | Display | 4.3-inch RGB LCD, Resolution 800*480 |
|  | Sampling Accuracy | 12-bit |
|  | Synchronization Mode | Internal Synchronization |
|  | Connectors | USB Interface (dual-purpose for charger input)3.5mm Stereo Headphone JackExternal Sensor Input Interface |
|  | Headphone Requirement | Minimum 8 Ohms |
|  | SD Card | Standard Configuration: 16G ~ 64G |
|  | Built-in Battery | 3.7V/10000mAh Lithium Battery |
|  | Operating Time | Approximately 12 hours |
|  | Charger | DC 5V/3A |
|  | Operating Temperature | -20 ~ 50℃ |
|  | Humidity | 20% ~ 85% Relative Humidity |
|  | Dimensions & Weight | 21010035(mm), 0.4KG (Host Unit) |

4. Product Configuration List

This product consists of a detection host, optional sensors, and related accessories. Detailed configuration is shown in Table 1:

Table 1

| Name | Quantity | Unit | Remarks |
| --- | --- | --- | --- |
| 1. Host | 1 | Unit | Built-in TEV and AA ultrasonic sensors |
| 2. AA1 External Non-contact Ultrasonic Sensor (Optional) | 1 | Piece | For PD detection of switchgears |
| 3. AE External Contact Ultrasonic Sensor (Optional) | 1 | Piece | For PD detection of transformers, GIS, motors, etc. |
| 4. AA2 External Focused Ultrasonic Sensor (Optional) | 1 | Piece | For PD detection of overhead lines |
| 5. HFCT High-Frequency Current Sensor (Optional) | 1 | Piece | For PD detection of cables |
| 6. UHF Ultra-High Frequency Sensor (Optional) | 1 | Piece | For PD detection of GIS |
| 7. Test Cable | 1 | Set |  |
| 8. Host Charger + Charging Cable | 1 | Set |  |
| 9. Headphone | 1 | Piece | For monitoring discharge sounds |
| 10. Coupling Agent (Optional) | 1 | Piece |  |
| 11. SD Card + Card Reader | 1 | Set |  |
| 12. User Manual | 1 | Copy |  |
| 13. Factory Inspection Report | 1 | Copy |  |
| 14. Certificate of Conformity | 1 | Piece |  |
| 15. Warranty Card | 1 | Piece |  |